The latest scale design uses one Load Cell Amplifier per Load Sensor (4 total).
These slides describe:
1) how to connect 1 load sensor to an amplifier
2) how to calibrate 4 independent load sensors
3) how to calculate weight and center of gravity

---
The Load Sensor Circuit

---
The Load Sensor circuit
Black
R1
R1 decreases slightly with weight
Red
R2 increases slightly with weight
R2
White

---
A Wheatstone Bridge
E+ (Vcc)
R1
R3
S+
S-
R2
R4
The Load Sensor
E- (GND)

---
An easier way to think about
a Wheatstone Bridge:
It is two voltage dividers
5V
5V
R1
R3
V1
V2
R2
R4
GND
GND
V1 = V2 only if 5V / (R1 + R2) * R2 = 5V / (R3 + R4) * R4. That is, R1:R2 as R3:R4

---
Using a Load Cell Amplifier
5V
5V
Tiny differences in V1-V2 are amplified.
A slight change in R1 and R2 causes a slight change in V1, which causes a change in (V1-V2) * 128.
That final value is digitized and reported by the HX711. That value is proportional to the change in R1 and R2, which is proportional to the weight on the Load Sensor.
R1
R3
V1
V2
R2
+
-
R4
GND
GND
(V1-V2) * 128

---
Choosing R3 and R4
Because the Load Sensor's R1 is closely matched to R2, we must choose R3 and R4 to be nearly equal to each other. If they aren't, there will be a large voltage difference between V1 and V2, and the amplified difference will be too large to digitize.
A pair of +/-0.1%, 1K resistors may vary by as much 2 ohms (1 + 1 ohm), so we buy several 0.1% 1K resistors and pick a pair of them that are very close to the same resistance. Buying 21 resistors will let us choose at least 1 pair that are within 0.1ohm of each other (it's like picking matching socks)
5V
5V
R1
R3
V1
V2
R2
R4
GND
GND
(V2 – V1) * 128

---
Calibrating 4 separate load sensors.
The V2 scale design calls for a separate amplifier for each load sensor, letting the Arduino see the value for each separate load sensor instead of mixing the 4 load sensors' output with a Load Sensor Combinator Board.

---
Calibrating one Load Sensor
The HX711 library assumes that the resistance of the load sensor
changes linearly with the weight applied to it.
Formula for a line: Y = M * X + B
X is Weight
Y is reported raw value from the amplifier.
Given two {X, Y} pairs, you can calculate M (SCALE) and B (OFFSET):
M = (X2 – X1) / Y2 – Y1)
B = Y1 – M * X1

---
Calibrating 4 Load Sensors:
put the weight at the center
of the rectangle
When a weight W is an equal distance from the 4 Load Sensors A B C and D, each Load Sensor sees ¼ of the weight.
Place a known weight in the center of the scale, and read the values from the Load Sensors A B C and D. Separately record X (W/4), and Y for each Load Sensor.
Repeat the process for at least one more weight.
Then calculate M (SCALE) and B (OFFSET) separately for each Load Sensor.
D
C
W
A
B

---
Measuring Weight and Center of Gravity
from the readings from the 4 Load Sensors

---
Background:
The statics of “Class 1 levers”
The weight on the two support points must sum to the weight placed anywhere between them.
W1
W2
W = W1 + W2
Also, as in a seesaw, a fulcrum with two weights, the Moments must sum to zero.
L1
L2
W1
W2
W1 * L1 = W2 * L2

---
Calculating weight W
and center of gravity {X,Y}
Given the weights read by the
Load Sensors A, B, C, and D
and knowing the scale dimensions
U and V, calculate X, Y, and W.
W = A + B + C + D
To calculate X and Y, we use a
property of the center of gravity:
that for any line going through the
center of gravity, the Moments on
one side of the line balance (equal)
those on the other side of the line.
D
C
V
Y
A
B
X
U

---
Calculating X
Imagine a knife-edged fulcrum is
under the line PQ. Because PQ goes
through the center of gravity, the
Moments perpendicular to that line
segment must sum to zero.
AX + DX = C(U-X) + B(U-X)
so
AX + BX + CX + DX = CU + BU
so
X = (C + B) / (A + B + C + D) * U
D
C
Q
P
A
B
X
U

---
Calculating Y
We calculate Y in the same way
we calculated X:
Imagine a knife-edged fulcrum is
under the line RS. Because RS goes
through the center of gravity, the
Moments perpendicular to that line
segment must sum to zero.
AY + BY = C(V-Y) + D(V-Y)
so
AY + BY + CY + DY = CU + DU
so
Y = (C + D) / (A + B + C + D) * V
D
C
V
R
S
Y
A
B

---
The Dog's center of gravity and weight
What we've measured so far is the weight and center of gravity of the combined plywood top plus Pippa (the dog).
To calculate the location of Pippa's center of gravity and her weight, we have to do one more lever calculation: Imagine a fulcrum at W
T * L1 = P * L2
W = T + P
We know T (plywood weight) from calibration.
We know L1 and W from our calculations from A, B, C, and D.
Find P (Pippa's weight) and L2.
So
P = W – T
L2 = T * L1 / P
From there we can use vector addition to calculate the X and Y of Pippa's center of gravity
D
C
T
P
L1
L2
A
B

---
Bad Math follows.
The following slides are my first attempt at calculating X and Y.
I kept this just to see where I went wrong.

---
Given W, L1, and L2, find W1 and W2
That is, if you place a known weight at a known position on a line, what are the weights at the two support points at the ends of that line?
W1
W2
L1
L2
We know W1 + W2 = W
So W2 = W – W1
We know W1 * L1 = W2 * L2
So W1 * L1 = (W – W1) * L2
Solving that for W1 gives: W1 = W * (L2 / (L1 + L2)
Check: When L1 = L2, W1 = W / 2: W * (L1 / 2 * L1) = W/2. Yes.
Check: When L1 = 0, W1 = W: W * (L2 / (0 + L2) = W * L2 / L2 = W. Yes.
Check: When L2 = 0, W1 = 0: W * (0 / (L1 + 0) = 0. Yes.

---
Dividing a weight among
4 support points on a rectangle
D
C
Known weight W
is at location {X, Y}
on a board of size {U, V}
A, B, C, and D are support points (load sensors).
Calculate the weight on each support point.
Q
V
Y
P
A
B
X
U

---
First calculate the weight at
points P & Q
Weight W is split among two points in a line. We know how to calculate the weights and the end points P and Q:
W1 = W * (L2 / (L1 + L2) and
W2 = W – W1
So
P = W * (V – Y) / V and
Q = W – P
That was easy!
Q (W2
L2 =
V - Y
V
Y (L1)
P (W1)

---
Given P, calculate A and B.
Given Q, calculate D and C.
Now that we know P, we can calculate A and B:
W1 = W * (L2 / (L1 + L2) and
W2 = W – W1
So
A = P * (U – X) / U and
B = P – A
We can similarly calculate D and C, given Q:
D = Q * (U – X) / U and
C = Q – D
D
C
Q
X
U
P (W)
A
(L1)
B (L2)
X (L1)
U (L1 + L2)

---
Calculating W from the Load Sensor values
With the 4 load sensors calibrated, we can read the weight at A, B, C, and D

---
Now A, B, C, D are known.
Calculate X, Y and W
D
C
We take the formulas for calibrating (finding A, B, C, and D from X, Y, and W) and solve them for our unknowns.
Q
V
Y
P
A
B
X
U

---
Calculate Q and X from D, C, and U,
and P and X from A, B, and U
D
C
D = Q * (U – X) / U and
C = Q – D
So
Q = C + D and
X = C * U / (C + D)
Similarly,
P = A + B and
X = B * U / (A + B)
Q
V
Y
P
A
B
X
U

---
Calculate Y and W
from P, Q, and V
D
C
W = P + Q
P * Y = Q * (V – Y)
Solve the second equation for Y:
Y = Q * V / (P + Q)
Q
V
Y
P
A
B
X
U

---
Error checking
As a side-effect of calculating P & Q, we have two independent calculations of X. The difference of these two values of X will give us a hint about how much error is in the measurement of A, B, C, and D.
Similarly, instead of calculating P & Q, we could pair the 4 Load Sensors the other way, to calculate R and S, Y, X, and W. The difference between these values of X, Y, W and the values we calculated first will give us an idea of how far out of calibration the scale is.
D
C
Q
V
R
S
Y
P
A
B
X
U